## Slide 1
CT Metal Artifact Reduction
via Image-Domain Learning
Lars Gjesteby, Hongming Shan, Qingsong Yang, Ge Wang
Biomedical Imaging Center, CBIS/BME/SoE
Rensselaer Polytechnic Institute
November 19, 2017

## Slide 2: Metal Artifacts
Metallic/dense objects cause severe metal artifacts
Existing metal artifact reduction (MAR) algorithms are imperfect
Relevant to medical imaging and airport scans of luggage
Smart techniques (machine learning/deep learning/CNN/ etc.) can improve MAR

## Slide 3: Metal Artifact Reduction
Metal objects corrupt CT images
  Beam-hardening, scatter, noise
1. W. Kalender et al., Radiology, 164(2), 1987.
2. M. Bal and L. Spies, Medical Physics, 33(8), 2006.
3. E. Meyer et al., Medical Physics, 37(10), 2010 (Referred to as NMAR)
4. J. Stayman et al., IEEE Trans. Med. Imag., 31(10), 2012.
5. O. Watzke et al., European Radiology, 14(5), 2004.
6. C. Reft, et al., Medical Physics 30(6), 2003.
7. L. Gjesteby et al., IEEE Access, 4, 2016 (Literature review)
Projection completion1-3
Iterative reconstruction4
Image-based post-processing5
Image quality remains insufficient (RT planning6)
NMAR [3]

## Slide 4: Proton Therapy
Energy lost by charged particles is inversely proportional to the square of their velocity
Precise localization of tumor is essential to ensure successful treatment

## Slide 5
Data
Domain
Image
Domain

## Slide 6: Superiority Principle of Deep Learning
Train a convolutional neural network (CNN) on various types of artifacts
Take advantage of initial correction by proven MAR algorithms (NMAR)
Do no worse than the state-of-the-art
More inputs, better quality (conclusions from Prof. Yu’s group)
State-of-the-art MAR results
CNN
More artifact reduction;
Better image quality

## Slide 7: Data Generation
Voxelized phantoms from the Visible Human Project: pelvis and spine regions
Industrial-grade CT simulator (CatSim, GE Global Research Center)
Fan-beam geometry
120 kVp, 300 mA, 720 views
Metal-free (ground truth)
Titanium-added
  NMAR algorithm applied
60 image slices, 512x512
  100k 48x48 patches for training; 15k for testing/validation

## Slide 8: Learning in the Image Domain
Input image: Already with correction from NMAR algorithm
Output image: Further MAR through CNN
CNN: Five convolution and five deconvolution layers
Supervised learning: Ground truth from simulation and/or experiments
…
…
3x3 kernel size
Number of filters denoted by n
Batch normalization (BN) in the first three layers
ReLU for activation

## Slide 9: CNN with WGAN
VGG Network
Discriminator
Truth Feature
Output Feature
Perceptual Loss
Adversarial Loss
CNN Generator
Generated Output Image
Truth Image
Truth Image
TensorFlow
Hyperparameters:
  Learning rate:
  Batch size: 128
  # of epochs: 40
Loss
Tuning
Tuning

## Slide 10: Experiments
CNN-MSE
  MSE loss function
CNN-WGAN-VGG
  VGG perceptual loss in the WGAN framework
CPCA-WGAN-VGG
  VGG perceptual loss in the WGAN framework with a modified generator (contracting path convolutional autoencoder)

## Slide 11: Testing Loss Curves

## Slide 12
Metal-free Truth
Uncorrected
NMAR
CNN-MSE
CPCA-WGAN-VGG
CNN-WGAN-VGG
MSE: 0.0903
SSIM: 0.3590
PSNR: 10.4454
MSE: 0.0104
SSIM: 0.7085
PSNR: 19.8411
MSE: 0.0092
SSIM: 0.7228
PSNR: 20.3856
MSE: 0.0065
SSIM: 0.7600
PSNR: 21.8702
Spinal Fixation Rods

## Slide 13
Metal-free Truth
Uncorrected
NMAR
CNN-MSE
CPCA-WGAN-VGG
CNN-WGAN-VGG
Enlarged ROIs

## Slide 14
Metal-free Truth
Uncorrected
NMAR
CNN-MSE
CPCA-WGAN-VGG
CNN-WGAN-VGG
MSE: 0.1156
SSIM: 0.3431
PSNR: 9.3687
MSE: 0.0147
SSIM: 0.6357
PSNR: 18.3168
MSE: 0.0149
SSIM: 0.6371
PSNR: 18.2712
MSE: 0.0109
SSIM: 0.6783
PSNR: 19.6452
Hip Prostheses

## Slide 15
Metal-free Truth
Uncorrected
NMAR
CNN-MSE
CPCA-WGAN-VGG
CNN-WGAN-VGG
Enlarged ROIs

## Slide 16: MAR-CNN Training with Pseudo Truth

## Slide 17: Pilot Results (IP Pending at RPI)
Metal-free Truth
NMAR (CNN Input)
CNN-MSE
Uncorrected
CNN-WGAN-VGG

## Slide 18: Discussions & Conclusion
Deep learning helps recover anatomical structures and tissue types when existing MAR algorithms are insufficient
Image domain methods have merits
Pseudo-truth could be powerful
Data and image domains will be combined in the future

## Slide 19
Wenxiang Cong
Guang Li
Hongming Shan
Ruibin Feng
Tao Xu
Qingsong Yang
Matthew Getzin
Lars Gjesteby
Fenglei Fan
Qing Lyu
Thank You!
Biomedical Imaging Center
Principal Investigator: Ge Wang